[image: image1.jpg]


[image: image2.jpg]


宁波新芝药检科技有限公司

Ningbo Scientz Drug Control Technology Co.,Ltd

RQ810S Dissolution sampling system

Technical Parameters
[image: image3.jpg]



Ningbo Scientz Drug Control Technology Co., Ltd.

WWW.Scientzbio.com


Overview:

RQ810S is composed of RC810S dissolution tester connected to QY8 sampler. It has granule dosage, solvent temperature monitoring, automatic sampling point, and 30 seconds filtration capacity greater than 10ml. Its performance indicators are in full compliance with the requirements of the Chinese Pharmacopoeia (2015 edition) and the "Guidelines for Mechanical Calibration of Dissolution Testers".

Function parameter:

1. It can realize the test of basket method, paddle method, small cup method 1, dish method (large and small dish) 1, rotating tube method (long and short tube) 1, which satisfies the requirements of blank and control.

2. Machine head electric lift

3. Automatic high clutch

4. Automatic, independent drug administration

5. Two-color lighting system

6. Sampling frame electric lifting

7. Solvent alone detection

8. Stirring speed detection

9. Use a precision glass syringe as the sampling power.

10. Adopt high precision, high pressure, high chemical stability valve body.

11. Built-in micro printer.

12. Multi-user management.

13. WiFi, Bluetooth, RS232, RS485 expansion interface.

14. Biometric identification.

15. The internal use of high chemical stability of the Teflon pipeline.

16. Filter discharge function to reduce adsorption.

17. Has the same amount of rehydration function.

18. With separate piping systems to meet the blanking and standard control requirements of the dissolution test.

19. A cylindrical filter head having a pore size of 10 μm was sampled at the initial filtration, and a φ25 mm syringe filter having a pore size of 0.8 μm/0.45 μm/0.22 μm was selected as needed for the secondary filtration.

A 20.16-bit test tube rack can hold up to 128 test tubes.

21. The sampling tube is equipped with a silicone cap to prevent evaporation and contamination of the sample.

22. It can be operated online with the dissolution apparatus.

Technical indicators:

RC810S Dissolution tester

	Model
	RC810S

	Number of experimental devices
	8

	Water bath temperature
	RT—45℃

	Solvent temperature detection
	8 aisle

	Temperature resolution
	0.01℃

	Stirring speed
	25RPM—250RPM

	Stirring speed resolution
	0.1RPM

	Dosing channel
	8

	Head lift
	electric

	Sampling rack lifting method
	Electric, automatic

	Sample point resolution
	1ml

	Lighting system
	Red light, white light

	amount of users
	7 users

	user level
	Third level

	User program
	99

	Record saving
	128 *366 days


QY8 Sampler
	﻿Model
	QY8

	Sampling channel
	8

	Number of samples
	16

	Longest sampling interval
	99 hours 59 minutes

	Minimum sampling interval
	1 minute for the first time, the remaining 2-10 minutes (depending on the sample size, sampling method and whether it is rehydration)

	Continuous working time
	99 hours 59 minutes × 16

	Sampling filter cycle
	＜30 second

	Sample volume (or fluid volume)
	1ml—100ml

	Single tube distribution
	0ml—15ml

	Sampling accuracy
	RSD1%（10ml）

	User program
	99

	amount of users
	7 users

	user level
	Third level

	Record saving
	128 *366 days


Contact Number：0574-88350013            Address：No. 65, MuJin Road, National High-tech Zone, Ningbo
 Website：www.scientzbio.com         

